{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["Anticorps","Modif","Liste"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":1,"defined_names":2,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Anticorps","index":0,"n_rows":111,"n_cols":10,"rows_kept":111,"cut":false,"hidden":false,"merged":[[0,1,0,8]]}
[null,"Liste des anticorps",null,null,null,null,null,null,null,"R-IH-FOR-004"]
["Mis a jour le ","2017-10-18"]
["ANTICORPS","CLONE","Fournisseur","REF\n Fournisseur","REF \nLabo\n","DILUTION","Témoin positif","Divers","Tps Ac 1er","Fiche Technique"]
["34BE12","34BE12","Clinisciences/DBS","Mob 059-05",4200,"1/50","NA","NA","20'","mai -2011\n"]
["a ACTINE","1A4","Agilent","M0851",4402,"1/400","0.11 €               ","NA","20'","Actine"]
["a FP","a 1 Foetoprotein","Agilent","A0008",4404,"1/200","NA","NA","20'","AFP"]
["A103","A103  MELAN-A","Agilent","M7196",4400,"1/50","NA","Mélanomes       ","30’","A 103"]
["ACE","11.7","Agilent","M7072",4479,0.02,"NA","NA","20'","ACE"]
["AE1/AE3","AE1/AE3","Thermo","MS343-P",5708,"1/150","NA","NA","15'","2014-02-01"]
["ALC","2B11 et PD7/26","Agilent","M0701",4405,"1/100","NA","NA","20’","ALC"]
["ALK1","ALK1","Clinisciences/Santa Cruz","SC-53157",4202,"1/50","NA","PH 9 ","30'","2013-04-01"]
["ALK4","5A4","Clinisciences/DBS","Mob416-05",4203,"1/50","oui","PH9 ","20'","NA"]
["BCL1","SP4","Thermo","RH-9104-S",5002,"1/25","NA","PH 9 ","30'","2013-05-01"]
["BCL2",124,"Agilent","M0887",4407,"1/50","NA","NA","20'","BCL2"]
["BCL6","PG-B6p","Agilent","M7211",4408,"1/25","NA","PH 9","30'","BCL6"]
["Ber-EP4","Ber-EP4","Agilent","M0804",4409,"1/100","NA","NA","20'","Ber-EP4"]
["Béta Caténine",14,"Diagomics","BSB 5088",4524,"1/100","oui","NA","30'","NA"]
["BHCG","hCG","Agilent","A0231",4410,"1/1000","NA","NA","15’","BHCG"]
["CA 125","ov185:1","Clinisciences/DBS","Mob110-05",4204,"1/50","NA","NA","20'","2013-03-01"]
["CA 19-9","c241:5:1:4","Clinisciences/DBS","Mob109-05",4205,"1/25","NA","NA","20'","2013-05-01"]
["Calcitonine","CAL-3-F5","Agilent","M3509",4411,"1/1600","NA","NA","12'","NA"]
["Caldesmone","CALD-31","Diagomics","BSB6102",4512,"1/200","NA","NA","20'","2011-09-01"]
["Calrétinine","POLY","Invitrogen","18-0211",4700,"1/100","NA","NA","20'","2013-05-01"]
["CD 117","C KIT (Poly)","Agilent","A4502",4412,"1/50","NA","NA","20'","CD117"]
["CD 138","MI 15  Syndécan","Agilent","M7228",4413,"1/100","NA","NA","20'","CD138"]
["CD 19","CD19-163","Leica","NCL-L-CD19",4802,"1/50","NA","PH 9","30'","NA"]
["CD 7","LP15","Diagomics","BSB5166",4502,"1/50","NA","PH 9","30'","2015-02-01"]
["CD10","56C6","Leica/Novo",35345,4800,"1/50","NA","NA","30'","2014-12-01"]
["CD15","C3D-1","Agilent","M0733",4414,"1/50","NA","NA","30'","2013-03-01"]
["CD1a","O10","Agilent","M3571",4415,"1/50","NA","NA","20'","CD1a"]
["CD20","L26","Agilent","M0755",4416,"1/200","NA","PH 9","20'","2014-10-01"]
["CD23","SP23","Thermo","RM-9123-S",5703,"1/100","NA","NA","30'","2013-08-01"]
["CD3","SP7","Thermo","RM-9107-S",5704,"1/50","NA","NA","20'","2014-03-01"]
["CD30","Ber-H2","Agilent","M0751",4417,"1/100","NA","NA","15’","CD30"]
["CD31","JC70A","Agilent","M0823",4418,"1/50","NA","NA","30’","CD31"]
["CD34","QBEND10","Agilent","M7165",4419,"1/200","NA","NA","15’","CD34"]
["CD4","RBT-CD4","Diagomics","BSB5152",4501,"1/200","NA","NA","30'","2015-01-01"]
["CD43 ","DF-T1","Agilent","M0786",4421,"1/50","NA","NA","15’","CD43"]
["CD5","CD5/4C7","Agilent","M3641",4478,0.02,"           ","PH 9","30'","CD5"]
["CD56","123C3","Clinisciences","18-01525Mob-261)",4206,"1/50","NA","NA","30'","2014-01-01"]
["CD68","PG-M1","Agilent","N1576",4423,"prédilué","NA","NA","20'","CD68"]
["CD79 a","JCB117","Agilent","M7050",4424,0.02,"NA","NA","20'","CD79a"]
["CD8","C8/144B","Thermo","MS-457-S1",5705,"1/25","NA","NA","20'","2013-02-01"]
["CD99  MIC2","12 E7","Agilent","N1593",4425,"prédilué","NA","NA","20’","2013-02-01"]
["CdK4","DCS-31","Invitrogen/Biosource","AHZ0202",4701,"1/100","NA","NA","20'","2013-07-01"]
["CDX2","CDX2-88","Biocare /Eurobio ","CM226B",4601,0.02,"oui","PH9 ","40'","NA"]
["CHROMO A","DAK-A3","Agilent","M0869",4426,"1/200","NA","NA","15’","CHROMO A"]
["CK19","RCK 108","Agilent","M0888",4427,"1/50","NA","NA","20'","CK19"]
["CK20","KS 20.8","Agilent","M7019",4428,"1/100","NA","NA","20'","CK20"]
["CK5/6","D5/16 B4","Agilent","M7237",4430,"1/100","NA","PH 9","20'","CK5/6"]
["CK7","OV-TL 12/30","Agilent","M7018",4431,"1/100","NA","NA","15’","CK7"]
["CMV ","CCH2+DDG9","Agilent","M0854",4433,"1/50","oui","NA","20'","CMV"]
["c-Myc","Y69","Eurobio/Biocare","CME 415 CK",4602,"1/100","NA","NA","NA","NA"]
["DESMINE","D33","Agilent","M0760",4435,"1/100","NA","NA","20'","Desmine"]
["DOG1","DOG1.1","Clinisciences","Mob 466-05",4217,"1/400","oui","NA","NA","NA"]
["E CADHERINE","NCH 384","Agilent","M3612",4437,"1/100","NA","NA","20'","E-Cadherine"]
["EBV","LMPAb-1","Thermo","MS-1458-S1",5007,"1/25","NA","NA","20'","2013-03-01"]
["EMA","E29","Agilent","M0613",4438,"1/2000","NA","NA","15’","EMA"]
["ERB2","C-ERB-2 oncoprotein","Agilent","A0485",4444,"1/800","oui+/-","NA","30’","Cerb2"]
["FVIII","FactVIII","Agilent","A0082",4446,"1/200","NA","NA","20'","F VIII"]
["Glycophorine A","GA-R2","Diagomics","BSB5586",4503,"1/100","NA","PH 9","30'","2015-01-01"]
["Glypican-3","1G12","Diagomics/Zytomed","MSK067-05",4522,"1/300","oui","NA","NA","NA"]
["Hélico","Poly","DiagomicZytomed","RBK012",4504,"1/300","oui","NA","30'","2014-04-01"]
["Hépatocyte","OCH1E5","Agilent","M7158",4447,"1/50","NA","NA","20'","2014-06-01"]
["HHV8","13B10","Diagomics","BSB5649",4515,"1/50","NA","NA","20'","2013-01-01"]
["HMB45","HMB45","Agilent","M0634",4448,"1/50","NA","NA","30'","HMB45"]
["HMGA2-P1","Poly","Diagomics","59170AP",4518,"1/200","oui","PH9 ","NA","NA"]
["HPL","Poly","Diagomics/Zytomed","516-2984",4505,"1/100","NA","NA","20'","2013-08-01"]
["Inhibine","R1","  Clinisciences","MOB 435-05 Biosystems",4210,"1/50","NA","PH 9","20'","2013-05-01"]
["Kappa BOM","R10-21-F3","Agilent","A0191",4450,"1/10000","NA","pas de démasquage","10'","K BOM"]
["Kappa TISSU","R10-21-F3","Agilent","A0191",4450,"1/5000","NA","pas de démasquage","10'","K TISSU"]
["KI67","MiB-1","Agilent","M7240",4451,"1/200","NA","NA","30’","KI 67"]
["KL1","KL1","Clinisciences/DBS","Mob191-05",4211,"1/100","NA","NA","20'","2013-08-01"]
["Lambda","N10/2","Agilent","A0193",4453,"1/10000","NA","pas de démasquage","10'","Lambda"]
["Mammaglobine","31A5","Diagomics/Zytomed","RBK022-05",4523,"1/200","NA","NA",30,"NA"]
["MDM2","IF2","Invitrogen/Zymed","18-2403",4702,"1/100","NA","NA","20'","2012-04-01"]
["MLH1","G168-15","BD/Pharmingen",550838,4100,"1/20","NA","PH 9","30'","2013-03-01"]
["MPO","Poly","Agilent","A0398",4455,"1/2000","NA","NA","20'","MPO"]
["MSH2","FE11","Biocare/Eurobio","CM219CK",4603,0.02,"NA","PH 9",30,"2013-05-01"]
["MSH6","EP49","Diagomics","BSB6937",4513,0.01,"NA","PH 9","20'","2017-09-01"]
["MUC1","BSB-44","Diagomics","BSB6154",4507,"1/50","NA","NA","30'","NA"]
["MUM1","MUM1p","Agilent","M7259",4456,"1/25","NA","NA","20'","MUM1"]
["Myogénine","F5D","MM France","296M-14",5012,"1/50","NA","NA","30'","2013-03-01"]
["NAPSIN A","KCG1,1","Clinisciences","Mob 463-05",4213,"1/200","NA","NA","20'","2011-02-01"]
["NSE","BBS/NC.VI-H14","Agilent","N0873",4457,"1/50","NA","NA","20'","2013-03-01"]
["P 16","KIT","Roche","p16 CIN tec",5200,"prédilué","NA","PH 9","20'","2013-12-01"]
["P40","BC28","Diagomics/Zytomed","MSK097-05",4521,"1/400","NA","NA","NA","NA"]
["P504S      ","13H4","MMFrance/Thermo","RM-9103-A",5006,"1/100","NA","NA","20'","2014-06-01"]
["P53","DO7","Diagomics/BioSB","BSB5845",4508,"1/100","NA","NA","20'","2014-06-01"]
["P57","Kip2 Ab-6","Thermo","MS-1062-P1",5707,"1/800","NA","PH 9",20,"2014-03-01"]
["P63","4A4","Biocare /Eurobio","CM163A",4600,"1/100","NA","NA","30'","2015-07-01"]
["P63/504S","4A4 + 13H4 1,2,3","NA","NA","NA","NA","NA","NA","30’","NA"]
["Parvo B19","R92F6","Diagomics","BSB5857",4509,"1/50","NA","NA","20'","NA"]
["Pax5","BV6","Clinisciences ","RMAB027",4218,0.04,"NA","NA","30'","NA"]
["Pax8","poly rabbit","Diagomics","RBK047-05",4526,"1/100","oui","PH 9","30'","NA"]
["PDL1","22C3","Agilent","M3653",4483,0.01,"oui","PH9 ","40'","NA"]
["PLAP","8A9","Agilent","M7191-0,2ml",4459,"1/25","NA","NA","20'","PLAP"]
["PMS2","PMS2","BD Pharmingen",556415,4101,"1/150","NA","PH 9","30’","2018-03-01"]
["PS100","S100","Agilent","Z0311",4461,"1/1000","NA","NA","15'","PS100"]
["PSA","ER-PR8","Agilent","M0750",4462,"1/50","NA","NA","30'","PSA"]
["RE","EP1","Agilent","M3643",4463,"1/200","NA","PH 9","30’","RE"]
["RP","PGR-312","Leica",4801,4801,"1/200","NA","PH 9","20'","2014-09-01"]
["Stat-6","poly","Diagomics","BSB3221",4519,"1/400","oui","NA","NA","NA"]
["Sox 10","EP268","Diagomics","BSB2584",4529,"1/100","NA","PH 9","30'","NA"]
["Synaptophysine","poly","Agilent","A0010",4465,"1/100","NA","NA","30'","2014-11-01"]
["TdT","poly.","Diagomics/Bio SB","BSB5970",4510,"1/100","NA","NA","30'","2014-07-01"]
["TG6","DAK-Tg6","Agilent","N1565",4466,"prédilué","NA","NA","15’","2012-03-01"]
["TTF1 Cyto","SPT24","Leica","NCL-lL-TTF1",4803,"1/100","NA","NA","30'","NA"]
["TTF1 Histo","8G7G3/1","Clinisciences/DBS","Mob 285",4215,"1/100","NA","NA","30’","2014-02-01"]
["Vimentine","V9","Agilent","M0725",4472,"1/100 ","NA","NA","20'","Vimentine"]
["WT1","6F-H2","MM France/Cellmarque","F/348M",5015,"1/50","oui","NA","30'","2014-08-01"]

{"sheet":"Modif","index":1,"n_rows":1,"n_cols":12,"rows_kept":1,"cut":false,"hidden":false,"merged":[]}
["ANTICORPS","CLONE","Fournisseur","REF \nFournisseur","REF \nLabo\n","DILUTION","Témoin positif","Divers","Tps Ac 1er","Fiche Technique","Date de modification ","Médecin\n demandeur"]

{"sheet":"Liste","index":2,"n_rows":10,"n_cols":1,"rows_kept":10,"cut":false,"hidden":false,"merged":[]}
["Médecin"]
["A"]
["C"]
["E"]
["F"]
["G"]
["I"]
["J"]
["K"]
["L"]